Chlorine Monitor

Range Modification

Adjust the dip switches on the free Chlorine meter, CL3630, located on the bottom right hand corner (see picture above).

From Left to Right

- OFF

- ON

- OFF

- OFF

- OFF

- OFF

- ON

After adjusting the dip switches, calibrate the meter.

- Take a sample of the inline process water near the sensor.

- Find out the value of the sample using a DPD / titration test.  Please sure the pH of the sample is at 7.

- Adjust the trimmer marked “SENS.” on the chlorine meter located inside of the control panel to the sample value.  The electrode is now calibrated.